PLEASE NOTE: The following specification contains areas, highlighted with the [ ] symbol. In these areas, the engineer has to make a selection, add specific, project related information and has to delete what is not applicable for the specific project.
Online Ion Selective Chloride Analyser
1- or 2-channel online analyser for monitoring chloride. Allows for up to six sample streams by using sequencers.The analyser provides automatic, online monitoring based on ion selective technology analysis. The reactor is automatically cleaned after each cycle and the analyser performs for up to 30 days unattended. Free programmable auto calibration ensures a constant check of the quality of the measurement. The measuremtent method is based on volumetric analysis refering to a titrimetry. A reagent solution (titrant), whose chemical composition is tailored to the sample to be analysed, is continuously added to a constant volume of sample solution until the ensuing chemical reaction is completed.
Technical Data
Number of sample streams

1 or 2
Sampling mode
Fixed or loop
Measuring interval
5 - 999 min freely programmable
Sample temperature
5 - 50 °C Pressure range
0.5 - 6 bar Flow rate
40 - 300 L/h
Sample volume
200 - 1000 mL per cycle
Sampling quality
Maximum particles allowed in sample are 30 g/L with a maximum diameter of 2 mm
Power requirements
(Voltage)
Power requirements
(Hz)
Process connection
(drain)
Process connection
(air)

110 - 240 V AC
50 Hz
2 x cable glands for diameter 2-6 mm
3 x cable glands for diameter 13-18 mm
4/6 mm diameter tube connector
Weight
25 kg
Sensor
Chloride Ion Selective Electrode (ISE) type 8471B, reference electrode type 8483B
Temperature compensation

Pt100 temperature sensor
Ambient temperature
5 - 55 °C
Measuring range
Chloride 0.02 - 500 mg/L
Accuracy
± 2 - 4 %
Calibration method
Manual 2-point or 1-point calibration method
Current output
2 x 0/4 - 20 mA linear, logarithmic or bilinear programmable
Scope of delivery
[ ] (delete what is not needed) Analytical Instrument
- Range
[ ] Chloride Analyser, Low Range (0.02 - 5 mg/L Cl)
[ ] Chloride Analyser, Standard Range (0.5 - 500 mg/L Cl) [ ] Chloride Analyser, High Range (50 - 20000 mg/L Cl)
- Power Supply
[ ] 220 V / 50 Hz [ ] 240 V / 50 Hz [ ] 110 V / 50 Hz [ ] 110 V / 60 Hz
Accessories
[ ] Reagent canister holder (up to 4 x 10-liter canisters) [ ] Manual sample entry system
[ ] Spare parts kit (2 years), ISE analyser with chloride electrode
[ ] Fiberglass enclosure NEMA4X free-standing version (double compartment) [ ] Fiberglass enclosure NEMA4X version (single compartment)
[ ] Tropicalisation of the analyser
[ ] Fiberglass enclosure (IP54) for wall-mounting
[ ] Steel free standing cabinet for analyser
[ ] Second channel option for analyser ISE version
Multi Channel Options
[ ] 2 channels sequencer [ ] 3 channels sequencer [ ] 4 channels sequencer [ ] 5 channels sequencer [ ] 6 channels sequencer
Services
[ ] Manufacturer’s services to perform start-up on instrument to include basic operational training and certification of
performance of the instrument.
[ ] Manufacturer’s Service Agreement that covers all the manufacturer’s recommended preventative maintenance, regularly scheduled calibration and any necessary repairs beginning from the time of equipment start-up through to end user acceptance / plant turnover and the first 12 months of end-user operation post turnover.
Brand:
Hach
Product:
Polymetron 8810 Online Ion Selective Chloride Analyser
